## Slide 1
‘CARD’ – Characterization of Adjoint Response
for Ge Detectors
By G. Sjoden1, J. Fee2, B. O’Day2, J. Petrosky2
CTBTO Science and Technology Conference 2017, Vienna, Austria
Topic 3: ADVANCES IN SENSORS, NETWORKS AND PROCESSING—
Design of Sensor Systems and Advanced Sensor Technologies
1Air Force Technical Applications Center, Patrick Air Force Base, Florida
2Air Force Institute of Technology, Wright-Patterson Air Force Base, Ohio
Adjoint Transport
INTRODUCTION
ADJOINT EFFICIENCY MODEL
C/E Results
The views expressed in this poster are those of the authors and do not necessarily reflect the official policy or position of the Air Force, the U.S. Department of Defense, or the United States Government.
Gamma ray spectroscopy an important tool in ensuring treaty verification
Need to understand / characterize HPGe detector efficiency, e
e calculated via deterministic adjoint radiation transport analysis
Adjoint transport methodology produces high fidelity map of detector efficiency e for any photon energy
Adjoint mapping important to reveal subtle effects of source placement & sample geometry on detector e
Impacts overall sensitivity of detector measurement, detectable quantities being measured, etc.
Conclusions
EXPERIMENT
Correspondence author: glenn.sjoden@us.af.mil, +1 321 494 4402
AFIT HPGe Detector
Detector measurement compared to computed adjoint efficiency data provided consistent results which revealed a ~3% efficiency loss for electronic conversion relative to modeled results (“offset” value). Future work will focus on enhancing model fidelity (Sn mesh, quadrature refinement), and comparison with additional measurement data for better e corrections.
Forward Boltzmann Transport Eqn:
Multigroup approx.:
Forward multi-group transport operator H used to derive adjoint operator H+ using adjoint identity for real functions; < > used for phase space integration:
Adjoint Identity:
Leads to adjoint eqn:
Adjoint function computed using ‘adjoint source’ aliased to detector cross section leads to adjoint ‘detector efficiency’.
Response R in detector is computed from either forward: or adjoint:
e profiled with either many forward models (moving forward source to numerous new source locations, recomputing multigroup transport each time) OR run a SINGLE adjoint transport computation to map group adjoint e values aliased to detector response cross section as function of position, energy, etc.
Adjoint transport yields global phase space gamma efficiency e map for detector
Lead
Air
Ge Detector
Al
C
3-D Model Slice 0.375 MeV 0.625 MeV 1.375 MeV
Test Data for
reference source
Computational (S20/P1) results depict a dramatic change in efficiency with position and energy, as expected
Comparisons made for 0 cm on top of detector