## 0cm centered
Conducted on 20170501_T175__CAL.TKA which is the case of the MNC source directly above the center of the detector.
Oday
Energy [keV] | Lower x x1 | Lower y y1 | Upper x x1 #2 | Upper y y1#2 | linspace # | p1*x | p2 | Gross Counts | Continuum Counts | Net Counts | Net Counts Error | % Error | Observed Count Rate
60 | 188 | 44640 | 281 | 38730 | 94 | -63.55 | 56590 | 5101400 | 3918600 | 1182800 | 3003.3315 | 0.25392 | 13.6898
88 | 285 | 38490 | 410 | 35660 | 126 | -22.64 | 44940 | 8067700 | 4671100 | 3396600 | 3569.1456 | 0.10508 | 39.3125
122 | 419 | 35330 | 545 | 35700 | 127 | 2.937 | 34100 | 7686700 | 4510500 | 3176200 | 3492.449 | 0.10996 | 36.7616
159 | 606 | 32350 | 706 | 29830 | 101 | -25.2 | 47620 | 5545300 | 3140000 | 2405300 | 2947.0833 | 0.12252 | 27.8391
320 | 1304 | 27360 | 1362 | 27310 | 59 | -0.8621 | 28480 | 1745700 | 1612500 | 133200 | 1832.5392 | 1.3758 | 1.5417
392 | 1566 | 25360 | 1681 | 23290 | 116 | -18 | 53550 | 7292100 | 2821932 | 4470168 | 3180.2566 | 0.071144 | 51.7381
514 | 2082 | 19900 | 2195 | 16920 | 114 | -26.37 | 74810 | 5116700 | 2099600 | 3017100 | 2686.3172 | 0.089036 | 34.9201
662 | 2675 | 16780 | 2808 | 15830 | 134 | -7.143 | 35890 | 12903000 | 2185200 | 10717800 | 3884.3532 | 0.036242 | 124.0486
898 | 3666 | 13600 | 3795 | 14100 | 130 | 3.876 | -609.3 | 7188800 | 1800500 | 5388300 | 2998.2161 | 0.055643 | 62.3646
1173 | 4820 | 8384 | 4953 | 6406 | 134 | -14.87 | 80700 | 8744000 | 1077100 | 7666900 | 3133.8634 | 0.040875 | 88.7373
1333 | 5452 | 4454 | 5622 | 4041 | 171 | -2.429 | 17700 | 7938000 | 726860 | 7211140 | 2943.6134 | 0.04082 | 83.4623
1836 | 7532 | 2127 | 7727 | 1973 | 196 | -0.7897 | 8075 | 3911400 | 401800 | 3509600 | 2076.8245 | 0.059176 | 40.6204
T175 | 2016-11-01
Nuclide | Gamma Energy (keV) | Half-life (days) | Half-life (years) | Half-Life Uncertainty | Activity (mic-Ci) | Branch Ratio % | Gamma/s | Gamma/s Uncertainty +/- | Total Uncertainty % | Decay Const (1/days)
Am-241 | 60 |  | 432.17 | 0.66 | 0.02941 | 36 | 391.7412 | 12.144 | 3.1 | 4.3942e-06
Cd-109 | 88 | 462.6 |  | 0.7 | 0.2707 | 3.63 | 363.5772 | 11.2709 | 3.1 | 0.0014984
Co-57 | 122 | 271.79 |  | 0.09 | 0.01019 | 85.6 | 322.7377 | 10.3276 | 3.2 | 0.0025503
Te-123m | 159 | 119.7 |  | 0.1 | 0.01403 | 84 | 436.0524 | 13.9537 | 3.2 | 0.0057907
Cr-51 | 320 | 27.706 |  | 0.007 | 0.3389 | 9.86 | 1236.375 | 38.3276 | 3.1 | 0.025018
Sn-113 | 392 | 115.09 |  | 0.04 | 0.05109 | 64.9 | 1226.8242 | 38.0315 | 3.1 | 0.0060227
Sr-85 | 514 | 64.849 |  | 0.004 | 0.06171 | 98.4 | 2246.7377 | 69.6489 | 3.1 | 0.010689
Cs-137 | 662 |  | 30.17 | 0.16 | 0.04325 | 85.1 | 1361.8128 | 42.2162 | 3.1 | 6.2944e-05
Y-88 | 898 | 106.63 |  | 0.025 | 0.09622 | 94 | 3346.5316 | 103.7425 | 3.1 | 0.0065005
Co-60 | 1173 |  | 5.272 | 0.001 | 0.05101 | 99.86 | 1884.7277 | 58.4266 | 3.1 | 0.00036021
Co-60 | 1333 |  | 5.272 | 0.001 | 0.05101 | 99.98 | 1886.9925 | 58.4968 | 3.1 | 0.00036021
Y-88 | 1836 | 106.63 |  | 0.025 | 0.09622 | 99.4 | 3538.7792 | 109.7022 | 3.1 | 0.0065005
Test: |  | Calibration File (20170501_T175__CAL.TKA)
Start Date: | 2017-05-01
Live time: | 86400 |  | seconds
Age: | 181 |  | days
Nuclide | Activity (gamma/s) | Expected Counts | FEP Counts | Photon Energy | Arbitrary Efficiency
Am-241 | 391.4298 | 33819530.7238 | 1182800 | 60 | 0.034974
Cd-109 | 277.213 | 23951202.2661 | 3396600 | 88 | 0.14181
Co-57 | 203.4123 | 17574819.9382 | 3176200 | 122 | 0.18072
Te-123m | 152.8788 | 13208724.0725 | 2405300 | 159 | 0.1821
Cr-51 | 13.3523 | 1153641.3709 | 133200 | 320 | 0.11546
Sn-113 | 412.4372 | 35634571.064 | 4470168 | 392 | 0.12544
Sr-85 | 324.5994 | 28045390.6832 | 3017100 | 514 | 0.10758
Cs-137 | 1346.3857 | 116327728.2453 | 10717800 | 662 | 0.092135
Y-88 | 1031.832 | 89150287.3649 | 5388300 | 898 | 0.060441
Co-60 | 1765.7669 | 152562263.1794 | 7666900 | 1173 | 0.050254
Co-60 | 1767.8888 | 152745594.5592 | 7211140 | 1333 | 0.04721
Y-88 | 1091.1075 | 94271686.8518 | 3509600 | 1836 | 0.037229

## 0cm flush to side
ODAY_20170427_xxxx Conducted on 20170427_T175__Flush.TKA which is the source on top of the detector flush with the side of the detector
Oday
Energy [keV] | Lower x x1 | Lower y y1 | Upper x x1 #2 | Upper y y1#2 | linspace # | p1*x | p2 | Gross Counts | Continuum Counts | Net Counts | Net Counts Error | % Error | Observed Count Rate
60 | 188 | 30680 | 281 | 26640 | 94 | -43.44 | 38850 | 3443800 | 2694400 | 749400 | 2477.5391 | 0.3306 | 8.6736
88 | 285 | 26480 | 410 | 24510 | 126 | -15.76 | 30970 | 5270500 | 3212200 | 2058300 | 2912.5075 | 0.1415 | 23.8229
122 | 419 | 24550 | 545 | 24360 | 127 | -1.508 | 25180 | 4928700 | 3105500 | 1823200 | 2834.4664 | 0.15547 | 21.1019
159 | 606 | 22400 | 706 | 20760 | 101 | -16.4 | 32340 | 3580700 | 2179700 | 1401000 | 2400.0833 | 0.17131 | 16.2153
320 | 1304 | 19330 | 1362 | 19160 | 59 | -2.931 | 23150 | 1210200 | 1135300 | 74900 | 1531.5025 | 2.0447 | 0.8669
392 | 1566 | 17730 | 1681 | 16520 | 116 | -10.52 | 34210 | 4560400 | 1987200 | 2573200 | 2558.8279 | 0.099441 | 29.7824
514 | 2082 | 13530 | 2195 | 11890 | 114 | -14.51 | 43750 | 3273300 | 1450100 | 1823200 | 2173.3384 | 0.1192 | 21.1019
662 | 2675 | 12280 | 2808 | 11560 | 134 | -5.414 | 26760 | 7753300 | 1595900 | 6157400 | 3057.6462 | 0.049658 | 71.2662
898 | 3666 | 10280 | 3795 | 10830 | 130 | 4.264 | -5350 | 4927100 | 1372400 | 3554700 | 2509.8805 | 0.070607 | 41.1424
1173 | 4820 | 6025 | 4953 | 4123 | 134 | -14.3 | 74950 | 5749700 | 679790 | 5069910 | 2535.6439 | 0.050014 | 58.6795
1333 | 5452 | 2974 | 5622 | 2442 | 171 | -3.129 | 20040 | 5252500 | 464220 | 4788280 | 2390.9663 | 0.049934 | 55.4199
1836 | 7532 | 1114 | 7727 | 846.9 | 196 | -1.37 | 11430 | 2654400 | 191610 | 2462790 | 1687.0122 | 0.0685 | 28.5045
T175 | 2016-11-01
Nuclide | Gamma Energy (keV) | Half-life (days) | Half-life (years) | Half-Life Uncertainty | Activity (mic-Ci) | Branch Ratio % | Gamma/s | Gamma/s Uncertainty +/- | Total Uncertainty % | Decay Const (1/days)
Am-241 | 60 |  | 432.17 | 0.66 | 0.02941 | 36 | 391.7412 | 12.144 | 3.1 | 4.3942e-06
Cd-109 | 88 | 462.6 |  | 0.7 | 0.2707 | 3.63 | 363.5772 | 11.2709 | 3.1 | 0.0014984
Co-57 | 122 | 271.79 |  | 0.09 | 0.01019 | 85.6 | 322.7377 | 10.3276 | 3.2 | 0.0025503
Te-123m | 159 | 119.7 |  | 0.1 | 0.01403 | 84 | 436.0524 | 13.9537 | 3.2 | 0.0057907
Cr-51 | 320 | 27.706 |  | 0.007 | 0.3389 | 9.86 | 1236.375 | 38.3276 | 3.1 | 0.025018
Sn-113 | 392 | 115.09 |  | 0.04 | 0.05109 | 64.9 | 1226.8242 | 38.0315 | 3.1 | 0.0060227
Sr-85 | 514 | 64.849 |  | 0.004 | 0.06171 | 98.4 | 2246.7377 | 69.6489 | 3.1 | 0.010689
Cs-137 | 662 |  | 30.17 | 0.16 | 0.04325 | 85.1 | 1361.8128 | 42.2162 | 3.1 | 6.2944e-05
Y-88 | 898 | 106.63 |  | 0.025 | 0.09622 | 94 | 3346.5316 | 103.7425 | 3.1 | 0.0065005
Co-60 | 1173 |  | 5.272 | 0.001 | 0.05101 | 99.86 | 1884.7277 | 58.4266 | 3.1 | 0.00036021
Co-60 | 1333 |  | 5.272 | 0.001 | 0.05101 | 99.98 | 1886.9925 | 58.4968 | 3.1 | 0.00036021
Y-88 | 1836 | 106.63 |  | 0.025 | 0.09622 | 99.4 | 3538.7792 | 109.7022 | 3.1 | 0.0065005
Test: |  | Calibration File (20170426_T175_CAL_3__11.TKA)
Start Date: | 2017-04-26
Live time: | 86400 |  | seconds
Age: | 176 |  | days
Nuclide | Activity (gamma/s) | Expected Counts | FEP Counts | Photon Energy | Arbitrary Efficiency
Am-241 | 391.4384 | 33820273.7778 | 749400 | 60 | 0.022158
Cd-109 | 279.2976 | 24131315.2447 | 2058300 | 88 | 0.085296
Co-57 | 206.0227 | 17800360.5613 | 1823200 | 122 | 0.10242
Te-123m | 157.3698 | 13596753.38 | 1401000 | 159 | 0.10304
Cr-51 | 15.1315 | 1307364.2319 | 74900 | 320 | 0.057291
Sn-113 | 425.0459 | 36723964.6973 | 2573200 | 392 | 0.070069
Sr-85 | 342.419 | 29584999.5205 | 1823200 | 514 | 0.061626
Cs-137 | 1346.8095 | 116364344.9182 | 6157400 | 662 | 0.052915
Y-88 | 1065.9201 | 92095493.7709 | 3554700 | 898 | 0.038598
Co-60 | 1768.95 | 152837283.9376 | 5069910 | 1173 | 0.033172
Co-60 | 1771.0758 | 153020945.8049 | 4788280 | 1333 | 0.031292
Y-88 | 1127.1538 | 97386085.9662 | 2462790 | 1836 | 0.025289

## 7cm centered
ODAY_20170425_1715 Conducted on 20170425_T175_CAL_7cm_center.TKA which is the case of the MNC source 7cm directly above the center of the detector.
Oday
Energy [keV] | Lower x x1 | Lower y y1 | Upper x x1 #2 | Upper y y1#2 | linspace # | p1*x | p2 | Gross Counts | Continuum Counts | Net Counts | Net Counts Error | % Error | Observed Count Rate
60 | 188 | 6757 | 281 | 6087 | 94 | -7.204 | 8111 | 906970 | 603640 | 303330 | 1229.0688 | 0.40519 | 3.5108
88 | 285 | 5855 | 410 | 5615 | 126 | -1.92 | 6402 | 1347500 | 722580 | 624920 | 1438.7773 | 0.23023 | 7.2329
122 | 419 | 5798 | 545 | 5855 | 127 | 0.4524 | 5608 | 1212200 | 739910 | 472290 | 1397.1793 | 0.29583 | 5.4663
159 | 606 | 5358 | 706 | 5336 | 101 | -0.22 | 5491 | 906210 | 540010 | 366200 | 1202.5889 | 0.3284 | 4.2384
320 | 1304 | 4364 | 1362 | 4153 | 59 | -3.638 | 9108 | 274370 | 251250 | 23120 | 724.9966 | 3.1358 | 0.26759
392 | 1566 | 3754 | 1681 | 3457 | 116 | -2.583 | 7798 | 1058800 | 418120 | 640680 | 1215.286 | 0.18969 | 7.4153
514 | 2082 | 3048 | 2195 | 2649 | 114 | -3.531 | 10400 | 786900 | 324780 | 462120 | 1054.3624 | 0.22816 | 5.3486
662 | 2675 | 2714 | 2808 | 2702 | 134 | -0.09023 | 2955 | 1826200 | 362820 | 1463380 | 1479.5337 | 0.1011 | 16.9373
898 | 3666 | 2202 | 3795 | 2287 | 130 | 0.6589 | -213.6 | 1265300 | 291780 | 973520 | 1247.8301 | 0.12818 | 11.2676
1173 | 4820 | 1274 | 4953 | 743.3 | 134 | -3.99 | 20510 | 1502500 | 135720 | 1366780 | 1279.9297 | 0.093646 | 15.8192
1333 | 5452 | 527 | 5622 | 373.7 | 171 | -0.9018 | 5443 | 1363800 | 76904 | 1286896 | 1200.2933 | 0.09327 | 14.8946
1836 | 7532 | 104.4 | 7727 | 38.57 | 196 | -0.3376 | 2647 | 674180 | 13971 | 660209 | 829.5487 | 0.12565 | 7.6413
T175 | 2016-11-01
Nuclide | Gamma Energy (keV) | Half-life (days) | Half-life (years) | Half-Life Uncertainty | Activity (mic-Ci) | Branch Ratio % | Gamma/s | Gamma/s Uncertainty +/- | Total Uncertainty % | Decay Const (1/days)
Am-241 | 60 |  | 432.17 | 0.66 | 0.02941 | 36 | 391.7412 | 12.144 | 3.1 | 4.3942e-06
Cd-109 | 88 | 462.6 |  | 0.7 | 0.2707 | 3.63 | 363.5772 | 11.2709 | 3.1 | 0.0014984
Co-57 | 122 | 271.79 |  | 0.09 | 0.01019 | 85.6 | 322.7377 | 10.3276 | 3.2 | 0.0025503
Te-123m | 159 | 119.7 |  | 0.1 | 0.01403 | 84 | 436.0524 | 13.9537 | 3.2 | 0.0057907
Cr-51 | 320 | 27.706 |  | 0.007 | 0.3389 | 9.86 | 1236.375 | 38.3276 | 3.1 | 0.025018
Sn-113 | 392 | 115.09 |  | 0.04 | 0.05109 | 64.9 | 1226.8242 | 38.0315 | 3.1 | 0.0060227
Sr-85 | 514 | 64.849 |  | 0.004 | 0.06171 | 98.4 | 2246.7377 | 69.6489 | 3.1 | 0.010689
Cs-137 | 662 |  | 30.17 | 0.16 | 0.04325 | 85.1 | 1361.8128 | 42.2162 | 3.1 | 6.2944e-05
Y-88 | 898 | 106.63 |  | 0.025 | 0.09622 | 94 | 3346.5316 | 103.7425 | 3.1 | 0.0065005
Co-60 | 1173 |  | 5.272 | 0.001 | 0.05101 | 99.86 | 1884.7277 | 58.4266 | 3.1 | 0.00036021
Co-60 | 1333 |  | 5.272 | 0.001 | 0.05101 | 99.98 | 1886.9925 | 58.4968 | 3.1 | 0.00036021
Y-88 | 1836 | 106.63 |  | 0.025 | 0.09622 | 99.4 | 3538.7792 | 109.7022 | 3.1 | 0.0065005
Test: |  | Calibration File (20170404_T175_CAL.TKA)
Start Date: | 2017-04-25
Live time: | 86400 |  | seconds
Age: | 175 |  | days
Nuclide | Activity (gamma/s) | Expected Counts | FEP Counts | Photon Energy | Arbitrary Efficiency
Am-241 | 391.4401 | 33820422.3905 | 303330 | 60 | 0.0089688
Cd-109 | 279.7164 | 24167500.0495 | 624920 | 88 | 0.025858
Co-57 | 206.5488 | 17845814.8322 | 472290 | 122 | 0.026465
Te-123m | 158.2838 | 13675716.5497 | 366200 | 159 | 0.026777
Cr-51 | 15.5149 | 1340484.3696 | 23120 | 320 | 0.017247
Sn-113 | 427.6135 | 36945807.7798 | 640680 | 392 | 0.017341
Sr-85 | 346.0986 | 29902918.7658 | 462120 | 514 | 0.015454
Cs-137 | 1346.8943 | 116371669.6358 | 1463380 | 662 | 0.012575
Y-88 | 1072.8716 | 92696109.5817 | 973520 | 898 | 0.010502
Co-60 | 1769.5874 | 152892347.5535 | 1366780 | 1173 | 0.0089395
Co-60 | 1771.7138 | 153076075.5898 | 1286896 | 1333 | 0.0084069
Y-88 | 1134.5047 | 98021205.2385 | 660209 | 1836 | 0.0067354

## Offset Below
Conducted on20170428_T175_Down_33.TKA for which the source is flush with the side of the shielding and 33 cm below the top of the shielding (6cm below the detector face)
Oday
Energy [keV] | Lower x x1 | Lower y y1 | Upper x x1 #2 | Upper y y1#2 | linspace # | p1*x | p2 | Gross Counts | Continuum Counts | Net Counts | Net Counts Error | % Error | Observed Count Rate
60 | 188 | 5813 | 281 | 6042 | 94 | 2.462 | 5350 | 679060 | 557170 | 121890 | 1111.8588 | 0.91218 | 1.4108
88 | 285 | 6088 | 410 | 6128 | 126 | 0.32 | 5997 | 1123200 | 769630 | 353570 | 1375.8016 | 0.38912 | 4.0922
122 | 419 | 6161 | 545 | 6463 | 127 | 2.397 | 5157 | 1121900 | 801670 | 320230 | 1386.9283 | 0.4331 | 3.7064
159 | 606 | 6293 | 706 | 6696 | 101 | 4.03 | 3851 | 896420 | 655960 | 240460 | 1245.9454 | 0.51815 | 2.7831
320 | 1304 | 4182 | 1362 | 3961 | 59 | -3.81 | 9151 | 261950 | 240260 | 21690 | 708.6678 | 3.2673 | 0.25104
392 | 1566 | 3393 | 1681 | 3101 | 116 | -2.539 | 7369 | 865490 | 376640 | 488850 | 1114.5089 | 0.22799 | 5.658
514 | 2082 | 2669 | 2195 | 2288 | 114 | -3.372 | 9689 | 671100 | 282490 | 388610 | 976.5193 | 0.25129 | 4.4978
662 | 2675 | 2368 | 2808 | 2279 | 134 | -0.6692 | 4158 | 1489400 | 311330 | 1178070 | 1341.9128 | 0.11391 | 13.6351
898 | 3666 | 2022 | 3795 | 1969 | 130 | -0.4109 | 3528 | 1035900 | 259370 | 776530 | 1138.0993 | 0.14656 | 8.9876
1173 | 4820 | 1103 | 4953 | 743.2 | 134 | -2.705 | 14140 | 1249100 | 123550 | 1125550 | 1171.6015 | 0.10409 | 13.0272
1333 | 5452 | 489.2 | 5622 | 349.5 | 171 | -0.8218 | 4949 | 1140400 | 68177 | 1072223 | 1099.353 | 0.10253 | 12.41
1836 | 7532 | 107.1 | 7727 | 34.84 | 196 | -0.3706 | 2898 | 557660 | 13819 | 543841 | 755.9623 | 0.139 | 6.2945
T175 | 2016-11-01
Nuclide | Gamma Energy (keV) | Half-life (days) | Half-life (years) | Half-Life Uncertainty | Activity (mic-Ci) | Branch Ratio % | Gamma/s | Gamma/s Uncertainty +/- | Total Uncertainty % | Decay Const (1/days)
Am-241 | 60 |  | 432.17 | 0.66 | 0.02941 | 36 | 391.7412 | 12.144 | 3.1 | 4.3942e-06
Cd-109 | 88 | 462.6 |  | 0.7 | 0.2707 | 3.63 | 363.5772 | 11.2709 | 3.1 | 0.0014984
Co-57 | 122 | 271.79 |  | 0.09 | 0.01019 | 85.6 | 322.7377 | 10.3276 | 3.2 | 0.0025503
Te-123m | 159 | 119.7 |  | 0.1 | 0.01403 | 84 | 436.0524 | 13.9537 | 3.2 | 0.0057907
Cr-51 | 320 | 27.706 |  | 0.007 | 0.3389 | 9.86 | 1236.375 | 38.3276 | 3.1 | 0.025018
Sn-113 | 392 | 115.09 |  | 0.04 | 0.05109 | 64.9 | 1226.8242 | 38.0315 | 3.1 | 0.0060227
Sr-85 | 514 | 64.849 |  | 0.004 | 0.06171 | 98.4 | 2246.7377 | 69.6489 | 3.1 | 0.010689
Cs-137 | 662 |  | 30.17 | 0.16 | 0.04325 | 85.1 | 1361.8128 | 42.2162 | 3.1 | 6.2944e-05
Y-88 | 898 | 106.63 |  | 0.025 | 0.09622 | 94 | 3346.5316 | 103.7425 | 3.1 | 0.0065005
Co-60 | 1173 |  | 5.272 | 0.001 | 0.05101 | 99.86 | 1884.7277 | 58.4266 | 3.1 | 0.00036021
Co-60 | 1333 |  | 5.272 | 0.001 | 0.05101 | 99.98 | 1886.9925 | 58.4968 | 3.1 | 0.00036021
Y-88 | 1836 | 106.63 |  | 0.025 | 0.09622 | 99.4 | 3538.7792 | 109.7022 | 3.1 | 0.0065005
Test: |  | Calibration File (20170426_T175_CAL_3__11.TKA)
Start Date: | 2017-04-27
Live time: | 86400 |  | seconds
Age: | 177 |  | days
Nuclide | Activity (gamma/s) | Expected Counts | FEP Counts | Photon Energy | Arbitrary Efficiency
Am-241 | 391.4366 | 33820125.1657 | 121890 | 60 | 0.0036041
Cd-109 | 278.8795 | 24095184.6176 | 353570 | 88 | 0.014674
Co-57 | 205.4979 | 17755022.0649 | 320230 | 122 | 0.018036
Te-123m | 156.4612 | 13518246.1413 | 240460 | 159 | 0.017788
Cr-51 | 14.7577 | 1275062.413 | 21690 | 320 | 0.017011
Sn-113 | 422.4937 | 36503453.6835 | 488850 | 392 | 0.013392
Sr-85 | 338.7785 | 29270460.3014 | 388610 | 514 | 0.013277
Cs-137 | 1346.7248 | 116357020.6617 | 1178070 | 662 | 0.010125
Y-88 | 1059.0135 | 91498769.5943 | 776530 | 898 | 0.0084868
Co-60 | 1768.313 | 152782240.1526 | 1125550 | 1173 | 0.007367
Co-60 | 1770.4379 | 152965835.8748 | 1072223 | 1333 | 0.0070096
Y-88 | 1119.8505 | 96755081.8901 | 543841 | 1836 | 0.0056208

## Offset Above
ODAY_20170426_1330 Conducted on 20170426_T175_CAL_3__11.TKA which is the case of the bottom of MNC source 11 cm below the top of the shield and the center of the source is 3cm from the inner edge of the shielding.
Oday
Energy [keV] | Lower x x1 | Lower y y1 | Upper x x1 #2 | Upper y y1#2 | linspace # | p1*x | p2 | Gross Counts | Continuum Counts | Net Counts | Net Counts Error | % Error | Observed Count Rate
60 | 188 | 2018 | 281 | 1971 | 94 | -0.5054 | 2113 | 230310 | 187480 | 42830 | 646.3668 | 1.5091 | 0.49572
88 | 285 | 2026 | 410 | 2177 | 126 | 1.208 | 1682 | 366880 | 265820 | 101060 | 795.4244 | 0.78708 | 1.1697
122 | 419 | 2219 | 545 | 2258 | 127 | 0.3095 | 2089 | 377870 | 284250 | 93620 | 813.7076 | 0.86916 | 1.0836
159 | 606 | 2193 | 706 | 2259 | 101 | 0.66 | 1793 | 300980 | 224820 | 76160 | 725.1207 | 0.9521 | 0.88148
320 | 1304 | 1669 | 1362 | 1576 | 59 | -1.603 | 3760 | 102460 | 95769 | 6691 | 445.2292 | 6.6541 | 0.077442
392 | 1566 | 1317 | 1681 | 1167 | 116 | -1.304 | 3360 | 290280 | 144180 | 146100 | 659.1358 | 0.45115 | 1.691
514 | 2082 | 888.3 | 2195 | 731.9 | 114 | -1.384 | 3770 | 213900 | 92376 | 121524 | 553.4221 | 0.4554 | 1.4065
662 | 2675 | 721.8 | 2808 | 696.2 | 134 | -0.1925 | 1237 | 442870 | 95041 | 347829 | 733.4242 | 0.21086 | 4.0258
898 | 3666 | 569.5 | 3795 | 614.6 | 130 | 0.3496 | -712.2 | 315500 | 76958 | 238542 | 626.4647 | 0.26262 | 2.7609
1173 | 4820 | 309 | 4953 | 220.4 | 134 | -0.6662 | 3520 | 378020 | 35458 | 342562 | 643.0226 | 0.18771 | 3.9648
1333 | 5452 | 138.3 | 5622 | 85.61 | 171 | -0.3099 | 1828 | 345590 | 19166 | 326424 | 603.9503 | 0.18502 | 3.7781
1836 | 7532 | 24.09 | 7727 | 6.467 | 196 | -0.09037 | 704.8 | 170440 | 3003.1 | 167436.9 | 416.465 | 0.24873 | 1.9379
T175 | 2016-11-01
Nuclide | Gamma Energy (keV) | Half-life (days) | Half-life (years) | Half-Life Uncertainty | Activity (mic-Ci) | Branch Ratio % | Gamma/s | Gamma/s Uncertainty +/- | Total Uncertainty % | Decay Const (1/days)
Am-241 | 60 |  | 432.17 | 0.66 | 0.02941 | 36 | 391.7412 | 12.144 | 3.1 | 4.3942e-06
Cd-109 | 88 | 462.6 |  | 0.7 | 0.2707 | 3.63 | 363.5772 | 11.2709 | 3.1 | 0.0014984
Co-57 | 122 | 271.79 |  | 0.09 | 0.01019 | 85.6 | 322.7377 | 10.3276 | 3.2 | 0.0025503
Te-123m | 159 | 119.7 |  | 0.1 | 0.01403 | 84 | 436.0524 | 13.9537 | 3.2 | 0.0057907
Cr-51 | 320 | 27.706 |  | 0.007 | 0.3389 | 9.86 | 1236.375 | 38.3276 | 3.1 | 0.025018
Sn-113 | 392 | 115.09 |  | 0.04 | 0.05109 | 64.9 | 1226.8242 | 38.0315 | 3.1 | 0.0060227
Sr-85 | 514 | 64.849 |  | 0.004 | 0.06171 | 98.4 | 2246.7377 | 69.6489 | 3.1 | 0.010689
Cs-137 | 662 |  | 30.17 | 0.16 | 0.04325 | 85.1 | 1361.8128 | 42.2162 | 3.1 | 6.2944e-05
Y-88 | 898 | 106.63 |  | 0.025 | 0.09622 | 94 | 3346.5316 | 103.7425 | 3.1 | 0.0065005
Co-60 | 1173 |  | 5.272 | 0.001 | 0.05101 | 99.86 | 1884.7277 | 58.4266 | 3.1 | 0.00036021
Co-60 | 1333 |  | 5.272 | 0.001 | 0.05101 | 99.98 | 1886.9925 | 58.4968 | 3.1 | 0.00036021
Y-88 | 1836 | 106.63 |  | 0.025 | 0.09622 | 99.4 | 3538.7792 | 109.7022 | 3.1 | 0.0065005
Test: |  | Calibration File (20170426_T175_CAL_3__11.TKA)
Start Date: | 2017-04-26
Live time: | 86400 |  | seconds
Age: | 176 |  | days
Nuclide | Activity (gamma/s) | Expected Counts | FEP Counts | Photon Energy | Arbitrary Efficiency
Am-241 | 391.4384 | 33820273.7778 | 42830 | 60 | 0.0012664
Cd-109 | 279.2976 | 24131315.2447 | 101060 | 88 | 0.0041879
Co-57 | 206.0227 | 17800360.5613 | 93620 | 122 | 0.0052594
Te-123m | 157.3698 | 13596753.38 | 76160 | 159 | 0.0056013
Cr-51 | 15.1315 | 1307364.2319 | 6691 | 320 | 0.0051179
Sn-113 | 425.0459 | 36723964.6973 | 146100 | 392 | 0.0039783
Sr-85 | 342.419 | 29584999.5205 | 121524 | 514 | 0.0041076
Cs-137 | 1346.8095 | 116364344.9182 | 347829 | 662 | 0.0029891
Y-88 | 1065.9201 | 92095493.7709 | 238542 | 898 | 0.0025902
Co-60 | 1768.95 | 152837283.9376 | 342562 | 1173 | 0.0022414
Co-60 | 1771.0758 | 153020945.8049 | 326424 | 1333 | 0.0021332
Y-88 | 1127.1538 | 97386085.9662 | 167436.9 | 1836 | 0.0017193

## Efficiency Errors
| 2017-05-01 | 2017-04-26 | 2017-04-25 | 2017-04-28 | 2017-04-26
| Net Counts |  |  |  |  | Error
keV | 0cm C | 0cm F | 7cm C | O-below | O-above | 0cm C | 0cm F | 7cm C | O-below | O-above
60 | 1182800 | 749400 | 303330 | 121890 | 42830 | 3003.3315 | 2477.5391 | 1229.0688 | 1111.8588 | 646.3668
88 | 3396600 | 2058300 | 624920 | 353570 | 101060 | 3569.1456 | 2912.5075 | 1438.7773 | 1375.8016 | 795.4244
122 | 3176200 | 1823200 | 472290 | 320230 | 93620 | 3492.449 | 2834.4664 | 1397.1793 | 1386.9283 | 813.7076
159 | 2405300 | 1401000 | 366200 | 240460 | 76160 | 2947.0833 | 2400.0833 | 1202.5889 | 1245.9454 | 725.1207
320 | 133200 | 74900 | 23120 | 21690 | 6691 | 1832.5392 | 1531.5025 | 724.9966 | 708.6678 | 445.2292
392 | 4470168 | 2573200 | 640680 | 488850 | 146100 | 3180.2566 | 2558.8279 | 1215.286 | 1114.5089 | 659.1358
514 | 3017100 | 1823200 | 462120 | 388610 | 121524 | 2686.3172 | 2173.3384 | 1054.3624 | 976.5193 | 553.4221
662 | 10717800 | 6157400 | 1463380 | 1178070 | 347829 | 3884.3532 | 3057.6462 | 1479.5337 | 1341.9128 | 733.4242
898 | 5388300 | 3554700 | 973520 | 776530 | 238542 | 2998.2161 | 2509.8805 | 1247.8301 | 1138.0993 | 626.4647
1173 | 7666900 | 5069910 | 1366780 | 1125550 | 342562 | 3133.8634 | 2535.6439 | 1279.9297 | 1171.6015 | 643.0226
1333 | 7211140 | 4788280 | 1286896 | 1072223 | 326424 | 2943.6134 | 2390.9663 | 1200.2933 | 1099.353 | 603.9503
1836 | 3509600 | 2462790 | 660209 | 543841 | 167436.9 | 2076.8245 | 1687.0122 | 829.5487 | 755.9623 | 416.465
| Expected Counts |  |  |  |  | Error
keV | 0cm C | 0cm F | 7cm C | O-below | O-above | 0cm C | 0cm F | 7cm C | O-below | O-above
60 | 33819530.7238 | 33820273.7778 | 33820422.3905 | 33820125.1657 | 33820273.7778 | 1048405.4524 | 1048428.4871 | 1048433.0941 | 1048423.8801 | 1048428.4871
88 | 23951202.2661 | 24131315.2447 | 24167500.0495 | 24095184.6176 | 24131315.2447 | 742487.2702 | 748070.7726 | 749192.5015 | 746950.7231 | 748070.7726
122 | 17574819.9382 | 17800360.5613 | 17845814.8322 | 17755022.0649 | 17800360.5613 | 562394.238 | 569611.538 | 571066.0746 | 568160.7061 | 569611.538
159 | 13208724.0725 | 13596753.38 | 13675716.5497 | 13518246.1413 | 13596753.38 | 422679.1703 | 435096.1082 | 437622.9296 | 432583.8765 | 435096.1082
320 | 1153641.3709 | 1307364.2319 | 1340484.3696 | 1275062.413 | 1307364.2319 | 35762.8825 | 40528.2912 | 41555.0155 | 39526.9348 | 40528.2912
392 | 35634571.064 | 36723964.6973 | 36945807.7798 | 36503453.6835 | 36723964.6973 | 1104671.703 | 1138442.9056 | 1145320.0412 | 1131607.0642 | 1138442.9056
514 | 28045390.6832 | 29584999.5205 | 29902918.7658 | 29270460.3014 | 29584999.5205 | 869407.1112 | 917134.9851 | 926990.4817 | 907384.2693 | 917134.9851
662 | 116327728.2453 | 116364344.9182 | 116371669.6358 | 116357020.6617 | 116364344.9182 | 3606159.5756 | 3607294.6925 | 3607521.7587 | 3607067.6405 | 3607294.6925
898 | 89150287.3649 | 92095493.7709 | 92696109.5817 | 91498769.5943 | 92095493.7709 | 2763658.9083 | 2854960.3069 | 2873579.397 | 2836461.8574 | 2854960.3069
1173 | 152562263.1794 | 152837283.9376 | 152892347.5535 | 152782240.1526 | 152837283.9376 | 4729430.1586 | 4737955.8021 | 4739662.7742 | 4736249.4447 | 4737955.8021
1333 | 152745594.5592 | 153020945.8049 | 153076075.5898 | 152965835.8748 | 153020945.8049 | 4735113.4313 | 4743649.32 | 4745358.3433 | 4741940.9121 | 4743649.32
1836 | 94271686.8518 | 97386085.9662 | 98021205.2385 | 96755081.8901 | 97386085.9662 | 2922422.2924 | 3018968.665 | 3038657.3624 | 2999407.5386 | 3018968.665
| Total Efficieny |  |  |  |  | Error |  |  |  |  | % Error
keV | 0cm C | 0cm F | 7cm C | O-below | O-above | 0cm C | 0cm F | 7cm C | O-below | O-above | 0cm C | 0cm F | 7cm C | O-below | O-above
60 | 0.034974 | 0.022158 | 0.0089688 | 0.0036041 | 0.0012664 | 0.0010878 | 0.0006908 | 0.0002804 | 0.00011646 | 4.3663e-05 | 3.1104 | 3.1176 | 3.1264 | 3.2314 | 3.4478
88 | 0.14181 | 0.085296 | 0.025858 | 0.014674 | 0.0041879 | 0.0043987 | 0.0026469 | 0.0008038 | 0.00045846 | 0.00013394 | 3.1018 | 3.1032 | 3.1085 | 3.1243 | 3.1984
122 | 0.18072 | 0.10242 | 0.026465 | 0.018036 | 0.0052594 | 0.0057866 | 0.0032815 | 0.00085049 | 0.00058241 | 0.0001744 | 3.2019 | 3.2038 | 3.2136 | 3.2292 | 3.3159
159 | 0.1821 | 0.10304 | 0.026777 | 0.017788 | 0.0056013 | 0.0058314 | 0.003302 | 0.00086138 | 0.00057662 | 0.00018701 | 3.2023 | 3.2046 | 3.2168 | 3.2417 | 3.3386
320 | 0.11546 | 0.057291 | 0.017247 | 0.017011 | 0.0051179 | 0.0039159 | 0.0021276 | 0.00076052 | 0.00076615 | 0.0003757 | 3.3916 | 3.7136 | 4.4094 | 4.5039 | 7.3408
392 | 0.12544 | 0.070069 | 0.017341 | 0.013392 | 0.0039783 | 0.0038898 | 0.0021732 | 0.00053858 | 0.00041627 | 0.00012463 | 3.1008 | 3.1016 | 3.1058 | 3.1084 | 3.1327
514 | 0.10758 | 0.061626 | 0.015454 | 0.013277 | 0.0041076 | 0.0033363 | 0.0019118 | 0.00048037 | 0.00041292 | 0.0001287 | 3.1013 | 3.1023 | 3.1084 | 3.1102 | 3.1333
662 | 0.092135 | 0.052915 | 0.012575 | 0.010125 | 0.0029891 | 0.0028564 | 0.0016406 | 0.00039003 | 0.00031407 | 9.2877e-05 | 3.1002 | 3.1004 | 3.1016 | 3.1021 | 3.1072
898 | 0.060441 | 0.038598 | 0.010502 | 0.0084868 | 0.0025902 | 0.001874 | 0.0011968 | 0.00032585 | 0.00026338 | 8.0583e-05 | 3.1005 | 3.1008 | 3.1026 | 3.1035 | 3.1111
1173 | 0.050254 | 0.033172 | 0.0089395 | 0.007367 | 0.0022414 | 0.001558 | 0.0010285 | 0.00027725 | 0.00022851 | 6.9609e-05 | 3.1003 | 3.1004 | 3.1014 | 3.1017 | 3.1057
1333 | 0.04721 | 0.031292 | 0.0084069 | 0.0070096 | 0.0021332 | 0.0014636 | 0.00097017 | 0.00026073 | 0.00021742 | 6.6247e-05 | 3.1003 | 3.1004 | 3.1014 | 3.1017 | 3.1055
1836 | 0.037229 | 0.025289 | 0.0067354 | 0.0056208 | 0.0017193 | 0.0011543 | 0.00078415 | 0.00020897 | 0.00017442 | 5.347e-05 | 3.1006 | 3.1008 | 3.1025 | 3.1031 | 3.11
Note on G2: O'Day, Buckley E Lt Col USAF AETC AFIT/ENP: Calculated using error propagation from the gross counts and the continuum
Note on B4: O'Day, Buckley E Lt Col USAF AETC AFIT/ENP: Data are taken from the individual sheets.
Note on G17: O'Day, Buckley E Lt Col USAF AETC AFIT/ENP: Calculated from the uncertainty in the source activity
Note on G32: O'Day, Buckley E Lt Col USAF AETC AFIT/ENP: Error propagation accounting for the uncertainty in the source and in our counts
Note on L32: O'Day, Buckley E Lt Col USAF AETC AFIT/ENP: Basically our uncertainty is driven by our uncertainty in the source activity

## Efficiency
|  | 0cm centered |  |  | 0cm Flush |  |  | 7cm centered |  |  | Offset Above |  |  | Offset Below
Nuclide | Energy | Source # | Det # | Efficiency | Source # | Det # | Efficiency | Source # | Det # | Efficiency | Source # | Det # | Efficiency | Source # | Det # | Efficiency
Am-241 | 60 | 33819530.7238 | 1182800 | 3.4974 | 33820273.7778 | 749400 | 2.2158 | 33820422.3905 | 303330 | 0.89688 | 33820273.7778 | 42830 | 0.12664 | 33820125.1657 | 121890 | 0.36041
Cd-109 | 88 | 23951202.2661 | 3396600 | 14.1813 | 24131315.2447 | 2058300 | 8.5296 | 24167500.0495 | 624920 | 2.5858 | 24131315.2447 | 101060 | 0.41879 | 24095184.6176 | 353570 | 1.4674
Co-57 | 122 | 17574819.9382 | 3176200 | 18.0724 | 17800360.5613 | 1823200 | 10.2425 | 17845814.8322 | 472290 | 2.6465 | 17800360.5613 | 93620 | 0.52594 | 17755022.0649 | 320230 | 1.8036
Te-123m | 159 | 13208724.0725 | 2405300 | 18.2099 | 13596753.38 | 1401000 | 10.3039 | 13675716.5497 | 366200 | 2.6777 | 13596753.38 | 76160 | 0.56013 | 13518246.1413 | 240460 | 1.7788
Cr-51 | 320 | 1153641.3709 | 133200 | 11.546 | 1307364.2319 | 74900 | 5.7291 | 1340484.3696 | 23120 | 1.7247 | 1307364.2319 | 6691 | 0.51179 | 1275062.413 | 21690 | 1.7011
Sn-113 | 392 | 35634571.064 | 4470168 | 12.5445 | 36723964.6973 | 2573200 | 7.0069 | 36945807.7798 | 640680 | 1.7341 | 36723964.6973 | 146100 | 0.39783 | 36503453.6835 | 488850 | 1.3392
Sr-85 | 514 | 28045390.6832 | 3017100 | 10.7579 | 29584999.5205 | 1823200 | 6.1626 | 29902918.7658 | 462120 | 1.5454 | 29584999.5205 | 121524 | 0.41076 | 29270460.3014 | 388610 | 1.3277
Cs-137 | 662 | 116327728.2453 | 10717800 | 9.2135 | 116364344.9182 | 6157400 | 5.2915 | 116371669.6358 | 1463380 | 1.2575 | 116364344.9182 | 347829 | 0.29891 | 116357020.6617 | 1178070 | 1.0125
Y-88 | 898 | 89150287.3649 | 5388300 | 6.0441 | 92095493.7709 | 3554700 | 3.8598 | 92696109.5817 | 973520 | 1.0502 | 92095493.7709 | 238542 | 0.25902 | 91498769.5943 | 776530 | 0.84868
Co-60 | 1173 | 152562263.1794 | 7666900 | 5.0254 | 152837283.9376 | 5069910 | 3.3172 | 152892347.5535 | 1366780 | 0.89395 | 152837283.9376 | 342562 | 0.22414 | 152782240.1526 | 1125550 | 0.7367
Co-60 | 1333 | 152745594.5592 | 7211140 | 4.721 | 153020945.8049 | 4788280 | 3.1292 | 153076075.5898 | 1286896 | 0.84069 | 153020945.8049 | 326424 | 0.21332 | 152965835.8748 | 1072223 | 0.70096
Y-88 | 1836 | 94271686.8518 | 3509600 | 3.7229 | 97386085.9662 | 2462790 | 2.5289 | 98021205.2385 | 660209 | 0.67354 | 97386085.9662 | 167436.9 | 0.17193 | 96755081.8901 | 543841 | 0.56208
Energy | 0cm | 0cm Flush | 7cm center | Offset Above | Offset Below
60 | 3.4974 | 2.2158 | 0.89688 | 0.12664 | 0.36041
88 | 14.1813 | 8.5296 | 2.5858 | 0.41879 | 1.4674
122 | 18.0724 | 10.2425 | 2.6465 | 0.52594 | 1.8036
159 | 18.2099 | 10.3039 | 2.6777 | 0.56013 | 1.7788
320 | 11.546 | 5.7291 | 1.7247 | 0.51179 | 1.7011
392 | 12.5445 | 7.0069 | 1.7341 | 0.39783 | 1.3392
514 | 10.7579 | 6.1626 | 1.5454 | 0.41076 | 1.3277
662 | 9.2135 | 5.2915 | 1.2575 | 0.29891 | 1.0125
898 | 6.0441 | 3.8598 | 1.0502 | 0.25902 | 0.84868
1173 | 5.0254 | 3.3172 | 0.89395 | 0.22414 | 0.7367
1333 | 4.721 | 3.1292 | 0.84069 | 0.21332 | 0.70096
1836 | 3.7229 | 2.5289 | 0.67354 | 0.17193 | 0.56208
Rough Geometric Corrections
|  | 0cm centered |  |  | 0cm Flush |  |  | 7cm centered |  |  | 16cm up 10.95 cm over |  |  | 6cm down 13.45cm over
Nuclide | Energy | Source # | Det # | Efficiency | Source # | Det # | Efficiency | Source # | Det # | Efficiency | Source # | Det # | Efficiency | Source # | Det # | Efficiency
Am-241 | 60 | 16909765.3619 | 1188300 | 7.0273 | 8455068.4444 | 763700 | 9.0325 | 2971797.0644 | 313930 | 10.5636 | 1305535.4563 | 43530 | 3.3343 | 2733917.8646 | 43530 | 1.5922
Cd-109 | 88 | 11975601.133 | 3535354 | 29.5213 | 6032828.8112 | 2071500 | 34.3371 | 2123595.7633 | 645470 | 30.3951 | 931520.7756 | 104990 | 11.2708 | 1947782.7286 | 104990 | 5.3902
Co-57 | 122 | 8787409.9691 | 3309400 | 37.6607 | 4450090.1403 | 1807000 | 40.6059 | 1568109.9283 | 483740 | 30.8486 | 687132.2805 | 90510 | 13.1721 | 1435262.9321 | 90510 | 6.3062
Te-123m | 159 | 6604362.0363 | 2868800 | 43.438 | 3399188.345 | 1390400 | 40.9039 | 1201683.8177 | 367210 | 30.558 | 524863.9838 | 77090 | 14.6876 | 1092774.6258 | 77090 | 7.0545
Cr-51 | 320 | 576820.6854 | 328200 | 56.8981 | 326841.058 | 95400 | 29.1885 | 117788.2248 | 30110 | 25.5628 | 50467.0769 | 7490 | 14.8414 | 103072.2356 | 7490 | 7.2667
Sn-113 | 392 | 17817285.532 | 5344900 | 29.9984 | 9180991.1743 | 2571200 | 28.0057 | 3246424.3596 | 638500 | 19.6678 | 1417624.1838 | 143960 | 10.155 | 2950830.1242 | 143960 | 4.8786
Sr-85 | 514 | 14022695.3416 | 4063400 | 28.9773 | 7396249.8801 | 1810700 | 24.4813 | 2627566.4206 | 468170 | 17.8176 | 1142044.7423 | 125682 | 11.005 | 2366136.5513 | 125682 | 5.3117
Cs-137 | 662 | 58163864.1227 | 10791000 | 18.5528 | 29091086.2296 | 6161400 | 21.1797 | 10225566.7362 | 1474000 | 14.4148 | 4491914.4991 | 351530 | 7.8258 | 9405953.878 | 351530 | 3.7373
Y-88 | 898 | 44575143.6824 | 6480400 | 14.5381 | 23023873.4427 | 3566600 | 15.4909 | 8145197.6902 | 968770 | 11.8938 | 3555084.5412 | 237200 | 6.6721 | 7396487.1376 | 237200 | 3.2069
Co-60 | 1173 | 76281131.5897 | 7810900 | 10.2396 | 38209320.9844 | 5077360 | 13.2883 | 13434634.9783 | 1360830 | 10.1293 | 5899848.5507 | 343288 | 5.8186 | 12350459.7839 | 343288 | 2.7796
Co-60 | 1333 | 76372797.2796 | 7297630 | 9.5553 | 38255236.4512 | 4784775 | 12.5075 | 13450779.1421 | 1282235 | 9.5328 | 5906938.2946 | 324413 | 5.4921 | 12365301.1135 | 324413 | 2.6236
Y-88 | 1836 | 47135843.4259 | 4198730 | 8.9077 | 24346521.4916 | 2426760 | 9.9676 | 8613113.3021 | 656680 | 7.6242 | 3759312.8021 | 167199.1 | 4.4476 | 7821391.7179 | 167199.1 | 2.1377
Energy | 0cm | 0cm Flush | 7cm center | Offset Above | Offset Below
60 | 7.0273 | 9.0325 | 10.5636 | 3.3343 | 1.5922
88 | 29.5213 | 34.3371 | 30.3951 | 11.2708 | 5.3902
122 | 37.6607 | 40.6059 | 30.8486 | 13.1721 | 6.3062
159 | 43.438 | 40.9039 | 30.558 | 14.6876 | 7.0545
320 | 56.8981 | 29.1885 | 25.5628 | 14.8414 | 7.2667
392 | 29.9984 | 28.0057 | 19.6678 | 10.155 | 4.8786
514 | 28.9773 | 24.4813 | 17.8176 | 11.005 | 5.3117
662 | 18.5528 | 21.1797 | 14.4148 | 7.8258 | 3.7373
898 | 14.5381 | 15.4909 | 11.8938 | 6.6721 | 3.2069
1173 | 10.2396 | 13.2883 | 10.1293 | 5.8186 | 2.7796
1333 | 9.5553 | 12.5075 | 9.5328 | 5.4921 | 2.6236
1836 | 8.9077 | 9.9676 | 7.6242 | 4.4476 | 2.1377
Note on A1: O'Day, Buckley E Lt Col USAF AETC AFIT/ENP: This is just some "stuff" I worked up because I was curious. Nothing important or rigorous here.

## Background
Oday |  |  |  |  | This shows the gamma lines and peak channels I used to perform the energy calibration on the background spectrum.
Energy [keV] | peak channel | f(x)=p1*x + p2
186 | 769 | p1 = | 0.2377 | (0.237,0.2384)
511 | 2141 | p2 = | 1.903 | (-1.246,5.051)
609.3 | 2559
911.2 | 3835
1460.8 | 6134
1764.5 | 7415

## Geometric Data
Detector well diameter | 27.9 cm | 279 mm | diameter as measured at the top of the well
Detector well depth | 40.6 cm | 406 mm | from top of well to bottom of the well
Depth to detector | 27 cm | 270 mm | from top of well to top of detector casing
Height of detector | 13.6 cm | 136 mm | from top of detector casing to bottom of the well
Detector center from wall | 13.95 cm
Position of 3-11 run | 10.95 cm | from detector center
| 16 cm | 16 cm above detector face or 11cm below the top of the well depth
Position of 33 run | 13.45 cm | from detector center (source placed such that the edge just touches the side of the well)
| 6cm | 6cm below the detector face or 7.6 cm from the bottom of the well
Calibration Run: | Source placed centered, directly on top of the detector housing.
0 cm flush to the side: | Source placed directly on top of the detecotr housing, with the edge of the source planchetted flush against the edge of the detector
7 cm centered: | Source placed centered, with the bottom of the source 7cm directly above the center of the detector
Offset above (3-11 run): | Source placed with the bottom of the source 16 cm above the top of the detector housing and the center of the source 10.95 offset from the center of the detector
Offset below (33 run): | Source placed with the bottom of the source 6 cm below the top of the detector housing (7.6 above the bottom of the well) with the planchette just touching the side of the well